## Slide 1
PERSONAL RESEARCH INFRASTRUCTURE
Alloy Lab Database
STAGE 2
The Characterisation Layer
Validated pipelines · Multi-instrument integration · ML-ready schema
PostgreSQL 17
Docker
customtkinter
OpenCV
scipy
numpy
matplotlib
July – September 2026
Notes: Stage 2 of the Alloy Lab Database project — built July to September 2026. Three validated characterisation pipelines for XRD, VSM, and SEM instruments.

## Slide 2
THE PROBLEM
Experimental data is scattered by default
📁
Instrument files
XRD patterns, VSM loops, SEM images — on instrument computers, shared drives, personal laptops.
📊
Spreadsheet analysis
Results manually copied into Origin projects, Excel files, or text documents after each measurement.
🧠
Implicit connections
"Which XRD pattern belongs to which sample?" lives in someone's head — or a sticky note.
"Stage 1 built the data-capture layer. Stage 2 adds the analytical layer — validated characterisation pipelines, stored in a schema designed for machine learning."
Notes: The fundamental problem: data scattered across instruments, analysis in spreadsheets, connections implicit. Stage 2 solves this by building validated, automated pipelines that store results directly in a structured database.

## Slide 3
STAGE 2 AT A GLANCE
Three instruments. One validated pipeline each.
3
characterisation pipelines
16
database tables · fully populated
170
XRD peaks · fitted & validated
690
SEM images imported
XRD
Peak fitting
NIST LaB₆ validated · lattice param within 0.0001 Å of certified value
VSM
Multi-segment hysteresis
Hc, Mr, BHmax, α/β coefficients, ΔSm entropy change · 3 instrument types
SEM
Unified metadata + phase fraction
Zeiss / JEOL / Tescan (16-bit) · operator-confirmed threshold
Notes: Each pipeline was built on real lab data and validated against a known reference before moving to the next. No shared code between instruments — each has its own failure modes.

## Slide 4
PIPELINE · XRD
XRD
X-ray Diffraction — Peak fitting pipeline
✓ Validated against NIST SRM 660c (LaB₆) — lattice parameter within 0.0001 Å of certified value
Background subtraction and Kα2 stripping (Cu anode)
Neighbour-aware peak fitting — each peak uses local context, not isolated windows
R² quality score per peak — low-quality fits flagged, never silently averaged
Scherrer crystallite size and d-spacing extracted automatically
xrd_peaks table: one row per peak (long format)
xrd_features table: one row per sample, ML-ready wide format
📷 XRD pattern with fitted peaks
(standalone analyzer tool)
57–58
peaks per sample
0.938
mean R²
0.0001 Å
vs NIST ref
Notes: XRD pipeline built on xrd_analyzer_dev1.py — fully general, not specific to any phase. Validated against a certified NIST reference material, not just against itself. Results in two tables: xrd_peaks (one row per peak) and xrd_features (one row per sample, ML-ready).

## Slide 5
PIPELINE · VSM
VSM
Vibrating Sample Magnetometry — multi-segment pipeline
Instruments: PPMS VSM · MPMS3 · ACMS | 3 real bugs found and fixed during integration, including a NaN masking issue that silently hid a 60 K temperature jump
Instrument auto-detection (PPMS/MPMS3/ACMS) — different column conventions per instrument
Mass extraction from file header (primary) or filename (fallback, with confidence flag)
Automatic segmentation: MH loops / MT scans / idle / corrupted
Rolling-window classification — not naive point-to-point derivatives
Second-quadrant Hc/Mr via crossing-point interpolation on descending branch only
Temperature coefficients α(Hc) and β(Mr) fitted across multiple-temperature series
Isothermal entropy change ΔSm — with check for genuine isotherms vs bipolar loops
Demagnetising correction and BHmax for cuboid samples
📷 VSM hysteresis loop — multi-segment file, per-segment Hc/Mr, accept/reject
Notes: The VSM pipeline handles multi-segment files (a single file can contain an MH loop + MT scan + temperature series). The entropy change check is important: a multi-temperature file with full bipolar loops looks structurally similar to a multi-isotherm entropy-change file, but using the wrong math gives silently wrong numbers.

## Slide 6
PIPELINE · SEM
SEM
Scanning Electron Microscopy — unified metadata + phase fraction
Zeiss
Proprietary compressed binary TIFF tags (34118/34119)
Pixel size, mag, EHT, detector — parsed from binary string
JEOL
Sidecar .txt file, custom $KEY value format
Scale bar in pixels + real units → exact pixel calibration. Footer height from $CM_FULL_SIZE vs actual TIFF height
Tescan
Sidecar .hdr file, standard INI format
PixelSizeX directly in metres. 16-bit image data — handled throughout pipeline
Design decision — why operator judgment is required
Otsu thresholding worked well on one real JEOL image and failed in two different ways on others — texture confusion on a Tescan image (which turned out to be real magnetic domain contrast, confirmed by Kerr microscopy), and noise confusion on Zeiss. A pipeline that pretends to automate a judgment it cannot reliably make is worse than one honest about its limits. The tool suggests Otsu as a starting point; the operator adjusts and saves the confirmed result.
Notes: The three-manufacturer problem was discovered mid-Stage-2 when JEOL files returned no metadata from the Zeiss-only parser. Each manufacturer was investigated directly from real files, not from documentation. The design decision on operator judgment came from a specific real finding: what looked like a segmentation artifact was actually real magnetic domain contrast, confirmed by a second imaging technique.

## Slide 7
INTERACTIVE TOOLS
Operator in the loop — Save to DB
Three standalone tools, each launched from the Data Viewer with the file pre-loaded and auto-analysed. Operator reviews, adjusts, and saves confirmed results to the database. No manual transcription.
Data Viewer
→
Pre-loaded file
→
Adjust params
→
💾 Save to DB
XRD
XRD Analyzer
· Live prominence/distance controls
· Per-peak accept/reject checkboxes
· Mean R² recalculated from accepted peaks only
· Save replaces previous analysis for that sample
VSM
VSM MH Analyzer
· Full multi-segment file loaded automatically
· Adjustable branch-detection sensitivity
· Per-segment accept/reject
· Save updates Hc/Mr for confirmed segments
SEM
Phase Fraction
· Otsu threshold as starting suggestion only
· Live slider, optional phase invert
· Calibrated area readout in µm²
· Save stores operator-confirmed result
Notes: The standalone tools share the same customtkinter framework as the main app, so they can later share the same database connection rather than being disconnected subprocesses. The "operator in the loop" design is a deliberate choice: the tools do not auto-save Otsu or default parameters — they require a human confirmation step.

## Slide 8
DATABASE SCHEMA
16 tables — designed for machine learning from the start
alloy_lab (PostgreSQL 17)
├── samples composition · lineage · VEC/δ/ΔH
├── compositions wt% + at% · aimed/measured/ref
├── synthesis method · temp · atmosphere
├── characterization one row per measurement
├── properties numeric results per measurement
├── literature_checks MP · OQMD · Alexandria
│
├── xrd_peaks one row per fitted peak
├── xrd_features ML-ready wide format
│
├── vsm_files file-level metadata
├── vsm_segments MH/MT/idle/corrupted
├── vsm_mh_details Hc · Mr · BHmax · flag
├── vsm_mt_details + vsm_mt_candidates
└── vsm_temperature_coefficients α · β
Feature matrix (X)
· Element fractions wt% + at%
· XRD: peak positions, widths, R²
· Crystallite size, d-spacing
· Phase fraction (SEM)
Target properties (y)
· Hc, Mr, BHmax
· Temperature coefficients α, β
· Entropy change ΔSm
· Synthesis route
→ Stage 3 ML
Notes: The two-table XRD schema (long + wide) is the key ML-design decision: xrd_peaks stores everything for audit and custom queries; xrd_features gives a ready-to-use fixed-width row per sample for scikit-learn or PyTorch. Same pattern could be extended to VSM and SEM in Stage 3.

## Slide 9
RESULTS
Real data, real validation — anonymised samples
Sample
System
XRD peaks
Mean R²
VSM Hc
Composition
Sample A
Rare-earth permanent magnet
57
0.938
✓
wt% + at%
Sample B
Rare-earth permanent magnet
58
0.943
—
wt% + at%
Sample C
Rare-earth permanent magnet
55
0.937
—
wt% + at%
Sample D
La-Fe-Si magnetocaloric
—
—
—
wt% + at%
Sample E
Mn-Al-C magnet
—
—
—
wt% + at%
Sample F
Fe-Co-P (synthesis blocked)
—
—
—
wt% + at% + routes
Sample names anonymised pending publication. Alloy systems cover rare-earth permanent magnets, magnetocaloric materials, and novel transition-metal compositions.
Notes: Sample F is particularly interesting: the screening identified P as a vapor-loss risk during iron melting (physical block, not just a caution), and automatically saved the three alternative synthesis routes — mechanical alloying, powder sintering, diffusion bonding — into the synthesis table with the full reasoning as notes.

## Slide 10
INFRASTRUCTURE · SEPTEMBER 2026
Three machines. One canonical database.
On 3 September 2026, the database was migrated from a local Docker container to a dedicated home server. One session, zero data loss.
M4 Mac
Development · app · Pi hub
✅ Connected · 50 samples
ML Desktop
Ryzen 9900X · RTX 5060 Ti
Stage 3 ML training
✅ Connected · 50 samples
Home Server
AMD FX-4320 · Ubuntu Server
PostgreSQL 17 · Docker
Wake-on-LAN enabled
→
→
50
samples
16
tables
170
XRD peaks
92
char. records
📷 Three terminal windows — Mac · ML Desktop · server — all confirming Samples: 50
Notes: Wake-on-LAN configured: server boots from cold in ~60 seconds via a magic packet. The ML desktop SSH was set up during this session. path_repair.py was run after migration and reported zero broken paths.

## Slide 11
NEXT · STAGE 3
The database is ready to train on
Everything built in Stage 2 was designed with this loop in mind:
🗄️
Feature matrix
compositions + structural features
→
🤖
GP / Neural Network
regression model
→
🎯
Property prediction
+ optimisation
→
🔬
New composition
recommendation
→
⚗️
New sample
→ measurements
What is already in the database
compositions
→ Element fractions wt% + at% · aimed / measured / reference
xrd_features
→ n_peaks · mean R² · crystallite size · d-spacing (ML-ready)
vsm_mh_details
→ Hc · Mr · BHmax per segment · quality flags
vsm_temp_coeff
→ α(Hc) · β(Mr) — temperature stability coefficients
properties
→ SEM phase fraction · pixel calibration
Notes: Stage 3 will run on the ML desktop (RTX 5060 Ti 16GB), reading from the server database. The plan is to start with Gaussian Process regression (good for small datasets) and move to neural networks as data grows. The closed loop — prediction → new sample → measurement → back into DB — is the goal.

## Slide 12
STAGE 2 · COMPLETE
The hard, instrument-specific work is done.
What remains is science.
✅ XRD pipeline — NIST validated
✅ VSM pipeline — 3 instrument types
✅ SEM pipeline — Zeiss / JEOL / Tescan
✅ Interactive tools with Save to DB
✅ 16-table ML-ready schema
✅ Element-fraction auto-wiring
✅ Synthesis route auto-save
✅ Database on home server
✅ Three machines connected
✅ path_repair.py — clean migration
What started as a way to stop losing track of lab notebooks is turning into a research platform.
github.com/iliya-radulov/my_db_project
Notes: Stage 3 starts with a clean infrastructure: validated data, structured schema, dedicated ML hardware, and a server that stays on and shared. The loop from composition → measurement → ML prediction → new composition is now possible to close.